# Controls Narrative

System Name: Pencil Module

Date: 07/25/2025

Author(s): M.Woodiwiss

# INTRODUCTION

The Pencil Module Filtration Test System is designed to automate and log performance data for pencil module filters. It enables precise control and monitoring of filtration, backwash, cleaning, and benchmark test cycles, ensuring repeatable and consistent test sequences. The system integrates pressure and temperature sensing, weight measurement, and both manual and automatic operation modes via an HMI interface.

# PROCESS OBJECTIVE

The system performs automated filtration, backwash, cleaning, and benchmark test cycles on pencil module filters. Throughout these processes, it continuously logs:

- Influent temperature

- Backwash supply pressure

- Influent supply pressure

- Weight of collected effluent and backwash water

- Operator-entered setpoints and process metadata

- Timestamps for each logged data point

The collected data is used to evaluate filter performance, cleaning effectiveness, and repeatability across multiple test cycles.

# INPUTS

| Signal | Type | Description |
| --- | --- | --- |
| Influent Supply Pressure | PIT | Monitors influent supply tank pressure |
| Backwash Supply Pressure | PIT | Monitors backwash supply tank pressure |
| Influent Temperature | RTD | Measures influent temperature |
| Effluent Weight | Scale | Measures filtered effluent |
| Backwash Weight | Scale | Measures backwash water output |
| Setpoints/Input Controls | HMI | Operator-configurable values (see Section 5) |
| Prime Button | HMI | Starts priming routine |
| Start/Stop Test Button | HMI | Begins/Stops automated cycle |
| Calibration Button | HMI | Applies offset to pressure/temp readings |

# OUTPUTS

| Device | Type | Function |
| --- | --- | --- |
| Influent Supply Valve | Solenoid | Controls influent supply flow |
| Backwash Supply Valve | Solenoid | Controls backwash supply flow |
| Effluent Valve | Solenoid | Controls effluent flow |
| Backwash Effluent Valve | Solenoid | Controls backwash effluent flow |
| Influent Water Drain Valve | Solenoid | Drains influent line during refill |
| Scale Zero Command | Digital | Zeros both scales at test start |

# SETPOINTS

- Filtration Target: Time (sec) or Weight (g)

- Backwash Target: Time (sec) or Weight (mL)

- Purge Time: Duration in seconds

- Cycle Count: Number of full cycles to perform

- Sample Time: Data logging interval, in seconds

- Calibration Offsets: Offsets used to calibrate sensors

- Project Name: Included in log file titles

- Sample ID: Included in log file titles

- Module ID: Included in log file titles

- FWD Soak Time: Duration in seconds

- BW Soak Time: Duration in seconds

# AUTOMATED PROCESSES

## Test

### Start Test:

- Zero both scales.

- Begin data logging at defined sample time intervals.

#### Cycle Loop:

##### Purge Phase

- Activate: Influent Supply Valve and Influent Water Drain Valve

- Run for [Refill Time] seconds

- Close all valves

##### Filter Phase

- Activate: Influent Supply Valve and Effluent Valve

- Direct water through filter to effluent collection

- Stop when either:

  - Effluent time = Effluent Time Setpoint, OR

  - Effluent weight = Effluent Weigh Setpoint

##### Backwash Phase

- Activate: Backwash Supply Valve and Backwash Effluent Valve

- Direct water through backwash line to backwash collection

- Stop when either:

  - Backwash time = Backwash Time Setpoint, OR

  - Backwash weight = Backwash Weight Setpoint

##### Cycle Completion

- Loop above until [Cycle Count] is reached

- Stop all outputs

- End data logging

- Return to idle state

## Clean

### Start Clean

- Zero both scales.

- Begin data logging at defined sample time intervals.

#### Cycle Loop:

##### Caustic:

###### Filtrate

- Activate: Influent Supply Valve and Effluent Valve

- Direct cleaning solution through filter to effluent collection

- Proceed when either:

  - Effluent time = Effluent Time Setpoint, OR

  - Effluent weight = Effluent Volume Setpoint

- Close valves

###### Backwash

- Activate: Backwash Supply Valve and Backwash Effluent Valve

- Direct cleaning solution backwash line to backwash collection

- Stop when either:

  - Backwash time = Backwash Time Setpoint, OR

  - Backwash weight = Backwash Volume Setpoint

- Close valves

###### Hold

- Soak solution for [Soak Time]

###### Filtrate

- Activate: Influent Supply Valve and Effluent Valve

- Direct cleaning solution through filter to effluent collection

- Proceed when either:

  - Effluent time = Effluent Time Setpoint, OR

  - Effluent weight = Effluent Volume Setpoint

- Close valves

###### Backwash

- Activate: Backwash Supply Valve and Backwash Effluent Valve

- Direct cleaning solution backwash line to backwash collection

- Stop when either:

  - Backwash time = Backwash Time Setpoint, OR

  - Backwash weight = Backwash Volume Setpoint

- Close valves

##### DI Rinse 1:

###### Filtrate

- Activate: Influent Supply Valve and Effluent Valve

- Direct cleaning solution through filter to effluent collection

- Proceed when either:

  - Effluent time = Effluent Time Setpoint, OR

  - Effluent weight = Effluent Volume Setpoint

- Close valves

###### Backwash

- Activate: Backwash Supply Valve and Backwash Effluent Valve

- Direct cleaning solution backwash line to backwash collection

- Stop when either:

  - Backwash time = Backwash Time Setpoint, OR

  - Backwash weight = Backwash Volume Setpoint

- Close valves

##### Acid:

###### Filtrate

- Activate: Influent Supply Valve and Effluent Valve

- Direct cleaning solution through filter to effluent collection

- Proceed when either:

  - Effluent time = Effluent Time Setpoint, OR

  - Effluent weight = Effluent Volume Setpoint

- Close valves

###### Backwash

- Activate: Backwash Supply Valve and Backwash Effluent Valve

- Direct cleaning solution backwash line to backwash collection

- Stop when either:

  - Backwash time = Backwash Time Setpoint, OR

  - Backwash weight = Backwash Volume Setpoint

- Close valves

###### Hold

- Soak solution for [Soak Time]

###### Filtrate

- Activate: Influent Supply Valve and Effluent Valve

- Direct cleaning solution through filter to effluent collection

- Proceed when either:

  - Effluent time = Effluent Time Setpoint, OR

  - Effluent weight = Effluent Volume Setpoint

- Close valves

###### Backwash

- Activate: Backwash Supply Valve and Backwash Effluent Valve

- Direct cleaning solution backwash line to backwash collection

- Stop when either:

  - Backwash time = Backwash Time Setpoint, OR

  - Backwash weight = Backwash Volume Setpoint

- Close valves

##### DI Rinse 2:

###### Filtrate

- Activate: Influent Supply Valve and Effluent Valve

- Direct cleaning solution through filter to effluent collection

- Proceed when either:

  - Effluent time = Effluent Time Setpoint, OR

  - Effluent weight = Effluent Volume Setpoint

- Close valves

###### Backwash

- Activate: Backwash Supply Valve and Backwash Effluent Valve

- Direct cleaning solution backwash line to backwash collection

- Stop when either:

  - Backwash time = Backwash Time Setpoint, OR

  - Backwash weight = Backwash Volume Setpoint

- Close valves

## Benchmark

### Start Benchmark

- Zero both scales.

- Begin data logging at defined sample time intervals.

#### Cycle Loop:

##### Purge Phase

- Activate: Influent Supply Valve and Influent Water Drain Valve

- Run for [Refill Time] seconds

- Close all valves

##### Filter Phase

- Activate: Influent Supply Valve and Effluent Valve

- Direct water through filter to effluent collection

- Stop when either:

  - Effluent time = Effluent Time Setpoint, OR

  - Effluent weight = Effluent Volume Setpoint

##### Backwash Phase

- Activate: Backwash Supply Valve and Backwash Effluent Valve

- Direct water through backwash line to backwash collection

- Stop when either:

  - Backwash time = Backwash Time Setpoint, OR

  - Backwash weight = Backwash Volume Setpoint

##### Cycle Completion

- Loop above until Repeat Count is reached

- Stop all outputs

- End data logging

- Return to idle state

# MANUAL OPERATION

A process flow diagram (PFD) is shown on the HMI.

Tapping any solenoid symbol on the HMI toggles that valve open/closed for manual testing or troubleshooting. Manual operation is only available when in idle mode

# CALIBRATION & SPECIAL FUNCTIONS

## Calibration

Users can apply calibration offsets to pressure and temperature readings via HMI.

## Scale Zeroing

Two individual manual zero buttons on the HMI, and an automatic scale zero command issued at the start of each test.